MULTIETHNIC STUDY OF ATHEROSCLEROSIS

COMPUTED TOMOGRAPHY

MANUAL OF OPERATIONS

EXAM III

January 1, 2005

I.  Background and Goals…………………………………………………………………………………..4

II.  Qualifications of Personnel	4

III.  Scanning Equipment and Centers	4

IV.  Training for Field Center CT Technologists and READERS	5

V.  Participant Scheduling	5

VI.  Scanning Table and Calibration Phantom.	6

VII.  Scanning Procedure	6

VIII. Scanner QualityAssurance………………………………………………………………………..10

IX.  Alerts	13

X.  Data Transmission 	13

XI.  Data Flow	13

XII.  CT Reading Procedure Operations Manual	23

INTRODUCTION	23

SUBJECT CHOOSER	24

1.  Launching The Subject Chooser	24

2.  The Subject Chooser Window	24

3.  Edit Scoring Defaults	25

BOXING	26

1.  Description Of Boxing	26

2.  Boxing The Heart	26

3.  Identifying The Regions In The Phantom	27

TRACING	27

1.  Launching The Tracing Program	27

2.  Labeling The Vessels	27

3.  Checking The Tracings	27

4.  Editing the X’s	27

5.  Finish Tracing	27

SCORING	28

1.  Launching The Scoring Program	28

2.  Checking Lesions	28

3.  Quality Assessment Tool	29

4. Auto Archiving (Automatic CD Backup)	30

APPENDIX A: Protocol for Inter-scan, Inter-reader, and Intra-reader Variability	31

1. Interscan Variability:	32

2. Inter-reader Variability:	32

3. Intra-reader Variability:	32

APPENDIX B: Description Of Options In Scoring Defaults Menu	32

APPENDIX C:  IDENTIFICATION AND REPORTING OF SCANNER PROTOCOL VIOLATIONS 34

APPENDIX C: CD Backup Procedure (Manual)	36

APPENDIX D: How To Make and Test A Test Case	37

XII.  Attachments	38

1.  MESa CT SCAN REPORT	39

2.  CT Study Transmission Sheet/ FAX to 310-533-1616	40

3.  QA QCT Sheet  FAX to 310-533-1616	41

4.  MESA CT Reading Center Complete Scan Log	.. 42

5.  MESA CT DATA FLOW SHEET.   ………………………………...……………………………………43                                                                                                          6.  MESA CT TECHNICIAN CERTIFICATION CHECKLIST…… .…………………….….……….....44

###### I.  Background and Goals

The Multiethnic Study of Atherosclerosis (MESA) is studying risk factors and measures of cardiovascular disease that relate to progression of subclinical to clinical disease.   An integral part of this study is the measurement of coronary artery calcium using either electron beam computed tomography (EBT) or multi-detector sub-second computed tomography (MDCT).   Coronary calcium will be assessed in relation to the risk of future cardiac events, and from repeated scans in selected individuals, the progression of coronary calcium will be related to baseline risk factors and risk of future events.

During exam III, approximately 3,350 participants will be scanned in the 6 centers between March 2004 and July 2005.

###### II. Qualifications of Personnel

Each Field Center has a designated physician to oversee the general performance of CT examination.  A reading center investigator will provide the necessary on-site training.

Field Center technologists should have appropriate knowledge of cross-sectional anatomy, physiology, and pathology of the heart.   Technologists must be certified as RTs in their state. It is recommended that technologists also have at least two years of experience in chest computed tomography.  The technologist should have a basic knowledge of cardiac CT, knowledge of computer software applications, data formatting, and experience with the workstations and data formatting / transmission procedures used.

Each technologist involved in the study should also have a complete understanding of this protocol, be experienced at providing breath-holding instructions, ECG gating and operation of the EBT or MDCT equipment. To ensure quality control, each Field Center should have designated CT technicians who will perform the MESA examinations.

###### III. Scanning Equipment and Centers

The scanners used will be:

Imatron C-150:			Saint Francis Hospital, Roslyn, NY

Imatron C-150:			Northwestern University, Chicago, IL

Imatron C-150:			UCLA School of Medicine, Los Angeles, CA

General Electric LightSpeed:		Wake Forest University, Winston-Salem, NC

Siemens S4+ Volume Zoom:		Johns Hopkins University, Baltimore, MD

Siemens S4+ Volume Zoom:		Univ. of Minnesota Medical Center, Minneapolis, MN

After Spring 2004:

Toshiba 32: 				Johns Hopkins University Medical Center, Baltimore, MD

Siemens Sensation 16:		Univ. of Minnesota Medical Center, Minneapolis, MN

General Electric LightSpeed Pro 16: 	Wake Forest University, Winston-Salem, NC

Each center will be responsible for scanning approximately one sixth of the cohort of 3350 individuals.  For purposes of increased reliability and quality control, each subject will be scanned twice during each session.

###### IV. Training for Field Center CT Technologists and Readers

CT technologists are already trained.   Site visits were done between September 11, 2002 and January 31, 2003.  At this visit, the site visitor witnessed at least two MESA participants being scanned.  Attention was paid to factors leading to mis-registration and motion artifact.  These are considered partially controllable by the technologists and are important for determining image quality and the accuracy of the data.

Additional or new technicians for MESA require certification and training either by a MESA CT investigator or lead technician who is already certified.   Certification requires completion of the checklist (Attachment 6), signed by the MESA CT lead technician performing the training, verifying that components of the scanning protocol have been reviewed with and understood by the new technician.   Then, three scans performed by the new technician should be sent to the MESA CT Reading Center with the ID# of each scan noted on the checklist in the spaces provided; once the CT Reading Center verifies that CT scanning have been done according to protocol, the new technician is considered fully certified to perform MESA CT scans and is so informed.   If certain aspects of the protocol are not adhered to, the technician may be asked to submit additional scans until these deficiencies are corrected.

A primary and back-up reader has been trained by the director of the reading center.  Periodic training sessions involving re-reading of cases for which there was discordance between reads of the same reader or of different readers are used for re-training.   Trained readers undergo a certification exam conducted by the reading center director.  The following essential elements  comprise this exam:

- Logging on and off of the work station

- Appropriate choosing of subjects

- Boxing and correcting the phantom regions

- Tracing of the major coronary arteries, including pruning of arterial tree

- Scoring and editing of lesions

- Editing when necessary the summary of scores

- Completing the quality assurance menu

The reader will be scored in each component above as either passed or failed.

###### V.  Participant Scheduling

The MESA operations committee will arrange details of scheduling. Each scanning site will provide their local Field Center with the days/times when MESA participants may be scheduled and the Field Center will contact participants to arrange their appointments for their CT scan.  Whether certain days/times will be reserved for MESA participants will be agreed upon by the Scanning Center and Field Center.

A Field Center interviewer/scheduler will be responsible for explaining and obtaining consent for the CT examination.  Participants will be scheduled and transportation arrangements, if necessary, will be arranged by the Field Center interviewer/scheduler.  A confirmation appointment letter will be sent, providing the time and date of appointment, directions to the scanning center and description of the procedure.  The CT completion form will be electronically sent to the Reading Center.  For Exam 3, this completion form also contains information on whether or not the subject has had angioplasty (with or without stent placement) or bypass surgery since the last exam.  This information is made available to the reader so that the reader can use it in the process of  reading scans.

###### VI. Scanning Table and Calibration Phantom

A special gel stuffed mat will be provided with the calibration pad phantom to prevent artifacts and for subject comfort.  This pad is placed between the phantom and the patient. The calibration phantom will be placed inside the blue catcher bag with its long axis parallel to the long axis of the scanning table.  The phantom is made of tissue equivalent plastic with rods of hydroxyapatite of known radiographic density. The technologist will position the subject so that the length of the phantom covers the expected length of the heart.

###### VII. Scanning Procedure

The scanning procedure described below, consisting of two scans done in succession on each subject, will require approximately 20 minutes of the subject’s time. This duration is required to provide minimal respiratory motion and maximum accuracy and reproducibility. In unusual cases, this may require as many as 30 minutes.  In many cases, the procedure will be completed in 15 minutes or less.

Entering the MESA ID and ACROSTIC.  This very important step must be done correctly to prevent any irretrievable loss of CT data.  All necessary information needs to be entered into the system including technician ID number and Field Center name.  The MESA seven digit ID number must be placed in the ID field and an acrostic as exemplified below will be placed in the name field.  The acrostic will consist of the first four letters of the last name followed by the first two letters of the first name, and then followed by one letter for the gender.  Each will be preceded by the letters “MESA”.

Example

ID:   MESA1212345		NAME:	MESADETRROM

The time stamp on each CT record will determine whether the scan is the first or second in the sequence of two scans.

Preparation of the Subject (3 minutes).  Subjects weighing more than 300 pounds will not be scanned due to technical difficulties.  The technologist will ask women if they might be pregnant and will not scan them if they answer affirmatively. The technologist will attach three electrocardiographic electrodes under the left clavicle and on either side of the thorax near the axillae (to maximize ECG signal). The technologist will ask each subject for their self reported weight in pounds and will use this value to decide if the subject satisfies the “Large Patient” criterion.

Breath Holding Instruction (3 minutes).  The technologist will instruct the subject on the importance of breath holding and immobility during scanning. An interpreter will assist in the instruction of subjects, not fluent in English. All centers will have access to interpreters fluent in languages spoken by at least 95%of potential volunteers.  Our preliminary study suggests that at least 99.5% will be able to hold their breaths for more than 15 seconds and 80% will be able to hold their breaths for more than thirty seconds.  Only after the technologist is satisfied that the subject understands the importance of breath holding, will he/she proceed.

Checking the Scout Image (1 minute).  The technologist will instruct the subject to take three deep breaths and then hold his/her breath at end-inspiration, then acquire an 11 cm scout image beginning 180 mm below the sternal notch. This will provide views of the chest on the image monitor at the operator console. From this, the technologist will check patient centering and choose the position for the highest scan (at the lower margin of the bifurcation of the main pulmonary artery).   The couch will be moved to the start position. The technologist will check subject and phantom positioning in the scout image.

Default Scanner Settings.

The technician will need to acquire at least 10.5 cm of data in the z direction for both of the scan series for MESA.  The 35 cm field of view will have to include the entire heart and the calibration phantom.  The most cephalad scan will be a the level of the pulmonary artery bifurcation and the most caudad scan will be beyond the apex of the heart.  All scans will use sequential axial mode.

Default Scanner Settings for Currently Used Scanners (as of 1/1/05)

Subjects weighing less than 220 pounds or 100 kg

- GE Light Speed Pro 16

- Gantry Rotation Time: 500msec.

- kVp: 120 kVp

- mA:  320 mA

- Exposure time: 326 msec

- Collimation:  4 x 2.5mm

- ECG Triggering:  50%

- Reconstruction.  Segmented

- Filter.  Standard/partial.

- Slice Thickness: 2.5 mm

- Reconstruction Field of View: 35 cm

- Siemens Sensation 16.

- Gantry Rotation Time:  500 msec.

- kVp:  140

- mAs:  50

- Quick Scan (Exposure time: 361 msec)

- Collimation: 12 x 1.5 mm

- ECG Trigger:  50%

- Reconstruction filter: B30F.

- Slice Thickness: 3mm

- Reconstruction Field of View: 35 cm.

- Siemens Sensation 64

Need to develop this

- Imatron EBT scanners (C150 or C300)

- kVp: 130 kVp,

- mA:  630

- Exposure Time: 100 msec.

- Collimation: 3mm

- ECG Trigger: 80%

- Reconstruction filter: Sharp

- Slice thickness:  3 mm

- Reconstruction Field of View: 35 cm

- Toshiba Aquillion 32:

  - Rotation time:  500 msec

  - kVp:  135

  - mA:	150

  - Exposure Time:  320 msec

  - Collimation:  4 x 3mm slices

  - ECG Trigger:  50%.

  - Use Medium filter with 320 mm scan field of view.

  - Reconstruction filter:   FC1 reconstruction kernel

  - Reconstruction Field of View:   32 cm

Large Patient Scanner Settings

- GE Light Speed Pro 16

- Gantry Rotation Time: 500msec.

- kVp: 120 kVp

- mA:  400 mA

- Exposure time: 326 msec

- Collimation:  4 x 2.5mm

- ECG Triggering:  50%

- Reconstruction.  Segmented

- Filter.  Standard/partial.

- Slice Thickness: 2.5 mm

- Reconstruction Field of View: 35 cm

- Siemens Sensation 16.

- Gantry Rotation Time:  500 msec.

- kVp:  140

- mAs:  63

- Quick Scan (Exposure time: 361 msec)

- Collimation: 12 x 1.5 mm

- ECG Trigger:  50%

- Reconstruction filter: B30F.

- Slice Thickness: 3mm

- Reconstruction Field of View: 35 cm.

- Siemens Sensation 64

Need to develop this

- Imatron EBT scanners (C150 or C300)

- kVp: 130 kVp,

- mA:  630

- Exposure Time: 100 msec.

- Collimation: 3mm

- ECG Trigger: 80%

- Reconstruction filter: Sharp

- Slice thickness:  3 mm

- Reconstruction Field of View: 35 cm

- Toshiba Aquillion 32:

  - Rotation time:  500 msec

  - kVp:  135

  - mA:	190

  - Exposure Time:  320 msec

  - Collimation:  4 x 3mm slices

  - ECG Trigger:  50%.

  - Use Medium filter with 350 mm scan field of view.

  - Reconstruction filter:   FC1 reconstruction kernel

  - Reconstruction Field of View:   35 cm

  - Slice Thickness:  3 mm.

Imaging (6 minutes).  All scanning will be done with a single breath hold.  The Siemens Volume Zoom helical scan acquisition requires a breath hold ranging from 15 to 25 seconds, depending upon the heart rate and heart length.  EBT and General Electric scan acquisition requires a breath hold of between 25 and 40 seconds depending on the same factors.  The technologist will always ask the subject to take two deep breaths and then a third before imaging commences.

Though total imaging time will be approximately 30 to 40 seconds, double scanning will require about 5 to 7 minutes to complete.  The technologist will first acquire one entire series of image slices. The technologist will instruct the subject to relax on the table while he/she reconstructs and assesses the adequacy of positioning, ECG gating and lack of respiratory motion. If these are adequate, the technologist will immediately acquire another series of image slices while the subject remains immobile and in an identical position.

Identification, Storage and Transport of Image Data.  The following will ensure proper scan identification and subject confidentiality. The technologist will record each volunteer’s study ID number in ID field of each scan. The acrostic will be put in the name field of the header. The date and times of the two scans and the scanning site name and technologist will be identified in the appropriate fields. The technologist will store images in retrievable format on optical disk (MODs).  Images will be transferred to a local work station which has an INTERNET interface to be used for transfer to the reading center.

###### VIII. Scanner Quality Assurance

###### Discussion of quality assurance is divided into three parts.  The first involves assurance that equivalent brightness levels are used and that brightness adjustment during reading is appropriate.   The second involves tracking of scan technique and quality and the third is concerned with assessment of reader accuracy and precision.

Brightness Levels.  Each clinical site will have a QCT calibration phantom and a torso QA phantom (Image Analysis Inc.) containing a central plug with a known concentration of hydroxyapatite (100mg/cc).  Quality assurance scans of the torso phantom allow convenient and quick verification of equivalence of brightness adjusted densities for the CT scanners at different sites.

## Procedure for Performing Brightness Calibration Checks

Adjust the table height so that the center of the calibration phantom is located at a distance of 9.2 cm +/- 1.0 cm from isocenter of the scanner field of view.  Place the torso phantom on top of the calibration phantom (positioned in couch pad); using the laser alignment light, adjust the table height until the torso center insert is at the isocenter (on grid).  This table height will be used for QA scans with the Torso phantom.

After the correct position has been determined, take a vertical axial slice through the center of the Torso phantom.  Use the same parameters as with patient exams.  With helical scanners, acquire two centimeters of volume data including the center of the torso phantom.    Reconstruction should be done with the same parameters as in scanning study subjects.

Using the axial image on the CT monitor and examine it to ensure that it is free of artifacts, such as air gaps and streaks.  Ensure that the calibration phantom is included in the field of view.  If there are significant artifacts over the calibration phantom, discard the image and rescan the phantom.

Using the CT software, place ROI’s on the calibration phantom reference samples (0, 50, 100, 200 mg/cc).  The 0 sample will be an apparent blank space on one end of the calibration phantom.  Then place an ROI in the TORSO vertebral sample.  The ROI’s should be as large as practical while remaining completely within the reference cylinder.  ROI’s about 70% of the sample area are recommended.  Record the five mean CT numbers within these five ROI’s using the standard QA sheet provided (Attachment 3). Record the five mean CT numbers within these five ROI’s.

File the data sheet for your records.  Send the QA QCT sheet by FAX (310-533-1616) to the study  reading center.

Instructions to the Reading Center

Use the available QCT software that runs on a standard PC using Windows 95.   Click on the QXCT-3000 icon on your desktop.  Click the enter QA button on the toolbar.  A QA data entry screen pops up.  Enter the data from the QA data sheets from the field centers.  When all the data is entered, click OK.  This will save the data to the QA database and open a window showing the QA report.  To discard the entries you have made, click cancel.

The QCT 3000 software computes the calibrated calcium density for the Torso phantom.   The results are displayed in graphic and tabular format.   The software also accesses the database and retrieves any previous data on the Torso phantom.  Previous data and calculated changes are displayed in the tabular form.

The individual QA torso readings should be maintained within + 3% of the mean value of all the readings.  If the values fall out of range, the field site must be notified in order to have the scanner checked by the field engineers.

The QA report has three major parts:

- Phantom and scanner information which includes the phantom type, exam number,  details of the CT scanner and technique used.

Click the image of the phantom which matches the one used at the field site and enter the 	serial number of the phantom.

- QA Readings, together with past readings, are plotted in graphical format.  The graph is color coded and includes a band that covers the mean + 3%.

- The QA readings are also printed in tabular form by date.  The change in the reading is also tabulated.  Also shown is the mean and standard deviation of all the readings.

Some common sources that cause poor results include inappropriate table height (Torso plug should be at isocenter when scanned), mis-positioning of the phantom, old or improper CT calibrations, or use of improper scan parameters.

When there is an unacceptable variability (greater than 3% from the mean), the reading center will inform the field site and an investigation into the possible causes of error will be undertaken. Further training/re-certification will be done if needed.  This can involve requesting that field engineers examine the scanning equipment.

Tracking of Scan Technique.  The coordinating center will review scans quarterly for inter-scan variability.  Cases of variability outside of the 98th percentile will be listed and the total brightness adjusted Agatston scores for these cases, identified by ID numbers, will be sent to the reading center.  The reading center director will review these cases to determine the potential for reader error. In cases of reader error, the director will re-read these cases in the presence of the original reader and re-transmit these results to be used in the corrected database.

Reader Assessments for Accuracy and Precision.  On a monthly basis, the co-director of the reading center will review a portion of scans read that week for accuracy of the arterial tracings.  Cases of questionable tracing will be reviewed with the director of the reading center and if the director concurs, he will address these inaccuracies with the CT reader responsible.

Quarterly, the coordinating center will generate a list of scans to be reread by the same reader and another list of scans to be reread by a different reader.  The results of these assessments will be relayed to the reading center director who will address cases of inter or intra reader discordance with the readers.  Scans will not be reread.

###### IX. Alerts

The Reading Center will review all scans within two weeks of receipt. If an obvious non-coronary abnormality is noted, this will be communicated to supervisory personnel immediately and a radiologist will review the scan. If the radiologist feels that an alert is warranted, he/she will write a report and this will be transmitted immediately to the field site PI, coordinator and a copy sent to the coordinating center.

###### X.  Data Transmission

The technologist will transmit the studies to the local work stations if this has not been done during scanning.  He/she will archive each study locally and leave a copy on the hard disk of the work station until it has been successfully transferred to the reading center.  He/she will transmit the studies electronically, using DICOM transfer on the local workstation, to the reading center on a given afternoon of each week.  He/she will simultaneously FAX the list of MESA ID numbers and acrostics to 310-533-1616.

The CT technical assistant will check each study to ascertain that it has been completely received.  If it has been so received, the study will be put on the queue to be read by the reader.   If it has not been received, the data manager will immediately communicate the identity of the study to the field site technologist and request that another attempt be made to send that study.   The CT data manager will enter the ID#, scan date, scan time and date and time of transfer of each study into a running log, a copy of which will be sent weekly to the coordinating center.

The CT reader will backup studies onto CDs as they are read.  The reader will label each disc with a volume name and the date the CD was created. The reader will print a directory to be stored with each CD.  Each CD-R that is archived will thus have a printed list of its contents inserted into its sleeve.  The Data Manager will check the CDs to be sure that they contain complete studies.

###### XI. Data Flow

###### Setting Up The Scan Log.  Data management will be done in collaboration and consultation with the coordinating center using a master spreadsheet.  The columns of this spreadsheet for exam III are shown below.   The reading center data management personnel will enter the scan completion date into the appropriate column from the forms e-mailed daily from the clinics to the reading center.

This scan log will reside on the snap servers in the Reading Center.

| ID # | Acrostic | Incorrect ID # Exam III | Date of Scan Exam III | Date Complete Scan Received Exam III | Number of Scans Received Exam III | Non- coronary Columns Exam III | Tech Asst Initials Exam III |
| --- | --- | --- | --- | --- | --- | --- | --- |

Interaction with Field and Scanning Site.

The data manager, supervisor and technical assistant will introduce themselves and get to know the field and scanning site personnel including the technologists, the field site coordinator, investigators, etc.

The field site coordinator will electronically transmit (e-mail, FTP, etc) the completed CT completion forms to the Reading Center on a weekly basis.  This form is completed by the coordinator and indicates the date when the CT scan was actually done.  The Technical Assistant will compile these forms daily and enter (cut and paste) the dates of scan completion in the third column of the scan log spreadsheet above.

The technologist will electronically transmit or mail CDs with CT scans to the reading center using DICOM transfer software on either end or using DICOM formatted CDs and overnight courier.  The technologist will FAX a transmission sheet (Attachment 2) to the reading center at the time of electronic transmission or mailing. This sheet will contain the ID numbers on the image studies of all scans transmitted or mailed.  The technical assistant at the reading center will collect the faxed transmission sheets (Attachment 2).  If there are any discrepancies between the transmission sheets and the studies actually received, these discrepancies will be settled between the technologist and technical assistant immediately so that the technologist can retransmit the scans.

###### Introduction to the Workstation.

###### When you first log on the workstation, the screen will look like the image at the right (Fig 11.1).  The blue part of the screen is called the Workspace.  All the programs you will be using will be shown on the Workspace in separate windows like Windows 98.  Moving around this operating system is just like moving around Windows.   You have a left mouse button to click on icons and commands.  To bring up various menus, use the right button.  Clicking the right button on any part of a window border will bring up various options.  These options will let you “CLOSE” the window, meaning minimize the window into a small icon, or “QUIT” programs and completely turn them off.

In the middle of the screen is a little gray box labeled “WorkSpace”, where you will launch various programs.  To bring up this “WorkSpace” box, you will click any part of the blue workspace using the third mouse button; this is known as “Right Clicking”.

Checking Receipt of Studies

The technical assistant will compare ID numbers on faxed transmittal sheets sent by the scan sites to the images processed by our workstation. To do this, the technical assistant will run the “que_man” program. Once this program has finished processing all the images, he/she will launch the “File Manager” to see the list of new images listed.  After sorting the files by date, he/she will check the studies off in the order they appear in the “File Manager.   He/she will recheck to be sure that all the studies on the sheets appear on the faxed transmission sheet.   Once a study is found to be complete (correct number of images/slices per scan, correct number of scans per patient), he/she will make a check mark on the faxed transmission sheet.    In the case of erroneous numbers, the technical assistant will copy the incorrect number in the column designated so that this can be corrected at the coordinating center.  After each study is checked for completeness, he/she will enter the dates that the complete scans were actually received at the reading center in the log sheet.   He/she will transmit this log to the coordinating center and the appropriate part of it to the respective field sites when he/she sends the summary of scores to the coordinating center.

He/she will inform the scanning site of all scans that are or are not received.  He/she will also e-mail a list of all received scans to the field site coordinators on a weekly basis.  The technical assistant will inform the scanning site about scans that need to be retransmitted.  The technical assistant will keep a binder with the faxed transmission sheets.

Processing DICOM Files. Log on as “read”. Begin the Image Processing step only when there are no incoming DICOM transmissions; we have noticed a tendency for the DICOM transmission to quit if Image Processing is carried out at the same time. On the top left corner of the workspace is a box with running graphs, Fig 11.2; this is the Performance Meter.  Each graph is running and shows what that particular component of the workstation is doing.

For now we are interested in the second box, labeled “pkts”. This graph displays the number of packets, or data, coming through the internet. Since we use the workstations solely for calcium scoring this line only shows activity if there are incoming electronic DICOM images or if someone else is logged on using X-Windows.  When this line is flat, as in Fig 11.2, it is okay to proceed with Image Processing. Otherwise wait until it is finished.

With the Faxed Transmission Sheets handy, type the following into the command tool window:

% cd /home/read/DICOM/CT

% que_man

“que_man” takes the raw DICOM files which have been stored in the “/home/read/DICOM/CT” directory and transfers them into the “/images” directory where they will be ready for the Readers to evaluate.  Once “que_man” is run, you will see various messages as it goes through each DICOM file.  During this period you will notice that the Hard-Drive is working and that the graphs in Fig 11.2 are active.

****	If “que_man” will not run, check the following:

****	% cd /home/read/DICOM/CT

****	% ls –lat | grep busy

****

****	If a file “.busy” is present, delete this file:

****	% rm –i .busy

****

****	Run “que_man” again, and it will now process the images normally.

Once “que_man” has finished processing all the images, launch the “File Manager”.  Using the “File Manager”, go to the “/images” directory and you will see all the new images listed.

After sorting the files by date, check the studies off in the order they appear in the “File Manager” (Fig 11.3) instead of the order on the faxed transmission sheet. Occasionally, a scanning technician will send a study and not include it in the faxed transmission sheet; if you will be checking off the faxed transmission sheet, you will miss these studies. Then recheck to be sure that all the studies on the sheets appear on the transmission sheet.

Double-click on a study folder in the “File Manager”; you will see a series of files labeled with long strings of numbers. These are the DICOM files and each one represents a single slice/image for an entire scan. The number of DICOM files in a study folder should be equal to the number of slices indicated in the faxed transmission sheet. Once a study is found to be complete (correct number of images/slices per scan, correct number of scans per patient), make a check mark on the faxed transmission sheet.  If there are any missing or incomplete studies, the data manager will have to contact the scanning site and inform them of the studies which need to be retransmitted.  In the case of incomplete studies, the entire study folder must be deleted to prevent a reader from inadvertently evaluating it.  To delete a study folder, right click on its folder and choose the delete option, Fig 11.4.

It is best to treat a faxed transmission sheet and the studies listed on it as a batch.  In this case, a batch is not complete unless all the studies listed on a transmission sheet have been checked off.  Treating the transmission sheets this way will make it easier to track which studies are needed for retransmission, and to address any questions a scanning site may have about a retransmission by looking at the sheet.  All unfinished transmission sheets should be in the front pocket of the study folder in case anyone needs to look at the information.

The Data Manager will then take the faxed transmission sheet and count the studies on it, and enter this number into its respective column in the cumulative study count log This log is located at \mydocument\chris\ebctscds\”Study Name” StatsMonthly.

Finally, the faxed transmission sheet will be placed in its appropriate place in the study folder.  Keep the faxed transmission sheets grouped together by scanning site and in chronological order.

Presentation to the Calcium CT Reader

The data manager will inform the CT reader of the new studies in the workstation.  It will not be necessary to enumerate the studies, since it will be the reader’s responsibility to read all NAMED studies. NAMED studies are scans that have been processed, but not yet read. Rather than providing a list of IDs (or the scan log) to the readers, this procedure keeps the study order randomized so that the reader does not read scan pairs in sequence (except rarely).

The Final Scan Log Sent to the Coordinating Center

As a final step, the technical assistant will submit the cumulative scan log on a weekly basis to the CT supervisor who will generate a report of cumulative work done and of outstanding scans not yet processed.  This will be performed every Friday.

###### Protocol Violations. During the boxing stage, the phantom finder software checks the image header of the scan for any deviation from its particular study protocol.  All protocols are listed in a text file in the “/home/read/bin” directory.  This is an ASCII text file containing relevant protocol information taken directly from the DICOM header of a particular scan site’s image.

If a protocol violation has been detected, a window (Fig 11.5) will pop up on the top left of the screen.  This window will list the relevant information in the DICOM header for the scan.  It is the reader’s responsibility to take note of these values and to inform the reading center supervisor of the violation.  The reading center supervisor will then check   the   header   information

against the proper protocols for that particular site and study.  If a violation has indeed occurred, the supervisor will inform the study coordinating center, the PI of the scanning site, and the lead technician of the scanning site.  The reader will then read the study normally.

###### House Keeping and Archiving.  It will be the Data Manager’s responsibility to periodically check the contents of the “/images”.  The number of studies in this folder should not exceed the maximum allowable, 248 image folders, by the Reading Software’s Subject Chooser, Fig 11.6.  When the number of image folders exceeds this number, the reader will only see 248 studies in the Subject Chooser and will not be able to read any of the excess studies.

If the number of studies in the subject chooser exceeds 248, there are relatively few options:

1. Move studies into a sub-directory temporarily.

2. Delete studies that have already been stored and archived.

Moving Files.  If the number of studies in the “/images” folder exceeds 248 and deleting is not an option, some studies will need to be temporarily moved to a subdirectory.  These studies should be moved in batches according to scan date, the earlier scanned studies to be left in the “/images” folder.  This procedure will ensure that the earlier studies are read first and that any inconvenience to the study coordinating center or to the scanning site will be kept to a minimum.

The best way to move studies in batches is through the “cmdtool”, or the Command Tool window.  This allows us to use UNIX wildcards and the study headers to move folders with ease.  First, determine the most recent studies that can be moved, noting their scan dates.

Then type in the following:

% cd /images

% mkdir studyname

% mv /images/*yyyymmdd* /images/studyname

Alternatively, this same procedure may be used to move batches of studies based on the study (for workstations with multiple studies) or scanning site.  Replace the “yyyymmdd” string in the third command into whatever string is common to all the studies you intend to move.

Deleting Scored/Archived Studies.  Studies should not be erased from the hard disk unless they have gone through the complete reading process, from boxing to scoring to archiving.  The safest way to delete studies from the hard disk would be to delete those that have already been archived onto CD’s.  Also, a benefit to using the CD’s as the basis for deleting is that by loading it onto the workstation and checking its directory structure, we inadvertently check if it went through the CD-burning process properly.

First, we will create a list of studies on a particular CD.  The lists I have previously created for the CARDIA and MESA studies are in “/home/read/temporary”.  For the MESA study, the lists are labeled as “mesalistxxx” where “xxx” is the cd number which the Reader created. For CARDIA, the lists are labeled “listxxx”.

Insert the CD in the cdrom drive of the workstation, please note the CD number.  Then type the following:

% cd /home/read/temporary

% ls /cdrom/cdrom0 > ./studylistxxx

You have created the file “studylistxxx” that contains an ascii-based list of the files on the CD.  Next, this file should be edited to reflect the actual folder names (see “Extra Info”) in the “/images” directory.  Open up the file you just created with “textedit”, the UNIX text based editor:

% textedit ./studylistxxx

Now, place a space before the last underscore of each entry (Fig 11.7) where the first five entries have been modified.   Notice that the last underscore separates the actual study folder name from the appended TAR creation date.  Now, save the file by right clicking on the “File” menu and choosing “Save”.  Then “Quit” textedit.

Run the following command which will strip the appended TAR creation date from entries in the file:

% more studylistxxx | cut –d” “ –f1 > temp

% mv temp studylistxxx

Now you have a list of study folder names formatted exactly like the study folder names in “/images”.  You can verify this by typing:

% more studylistxxx

The next step is to delete studies based on the entries on this list that were created.  Please note, type in the following command very carefully; any typo in any part of the command may delete EVERYTHING in the “/images” directory. Again, please type carefully, if anything, this should be done by someone who understands UNIX commands and who is comfortable running them.

% rm –r /images/`more /home/read/temporary/mesalistxxx’

After deleting, import the entries in “mesalistxxx” into an excel file.  For MESA and CARDIA, they are in an excel file (Fig11.8) called “MESA_Archive_MasterList” and “CARDIA_Archive_MasterList”, respectively.  This will help facilitate searching and reloading studies in the future.

Printing.  Printing files (and standard output) to the HPLJ printer

Operation

To print an ASCII file to the default printer in portrait mode, type:

% enscript filename

To print an ASCII file to the default printer in landscape mode, type:

% enscript filename –r

To print the text output of a UNIX command, type:

% command | enscript [-r]

Examples

ls | enscript

dcm_dump_file | enscript –r

Notes

Enscript is a GNU public domain program.  Additional -options [e.g. -G] are documented on the GNU web site.

Additional Archives.  The technical assistant will make two additional copies of each scan every week by copying all CD’s generated that week onto DVDs using the data management and archiving computer with high speed DVD and CD drives.   One CD will be mailed to the MESA coordinating center. The other will be kept on campus in another location.

Mailed CDs. Transferring DICOM data from CD-R’s

Operation

To transfer DICOM files from a CD-R, insert the CD into a network accessible PC running Windows 2000 or NT.  Double-click on the CD-TO-DET[1, 2, or 3] desktop icon.  Remove the disk when the DOS dialog box closes, and the multi-line output is no longer displayed.   Store the CD-R.   Transfer the resulting “big-endian” DICOM files on DET[1, 2, or 3] by invoking “que_man” on the UNIX workstation as normal.

Requirements

Three files must be present on the PC:

- CD-2-DET1,2,3].LNK  [A desktop shortcut, with target: C:\cd2det3_2k.bat]

- CD2DET[1,2,3]_2K.BAT  [A script which involves transfer of all DCM files on CD]

- CDR-TO-DCM.EXE  [A program which transfers data to det[1,2,3] via port 4000]

Notes

All DCM files on a CD will be transferred to the selected workstation unless the above Script file is run from the DOS-Prompt from C:\, after first selecting the appropriate parent directory of E:, and from which all DCM files will be selected.

On running this program, a constant stream of DOS-level messages should scroll. Each group parsed represents one slice transferred to the workstation.

IP address, port number [4000 is our default], and AE title may be changed by altering Default values in the CD2DET[1,2,3]_2K.BAT NT script file.

Reloading Studies. The files on the CD’s are TAR files, which are similar to ZIP files on PCs. One TAR file contains the whole study folder (including images and tracings).  It is also labeled with the same name as its original study folder in the “/images” directory, plus the date the TAR file was created.

The basic TAR file command is:

% tar xvf /cdrom/cdrom0/tarfilename

This command takes the TAR file on the CD and places all of the original files (DICOM files, tracings, and others) back into the proper directory in “/images”.  The above command is good for taking TAR files off of the CD one by one.

When reloading a large amount of studies from the CD’s it is best to make an ASCII based list and to use this list in a UNIX script.  Using the Excel Archive List file (Fig 11.8), it is easy to search for individual studies by ID number and by scan time.  Once a particular scan is found in this Excel file, highlight the whole row, copy and then paste into a new worksheet.  Once all studies that have to be reloaded are found, print the file. This print-out will be used to physically locate the CD on the shelves and to check off each study loaded.  Then take this list and save it as a tab-delimited text in the “/images” directory of the workstation where the studies are to be reloaded.

Next take the list you printed with the numbers of the CD’s needed and use this to find the CD’s on the shelf.  Next, place the CD in the workstation’s CD-drive and run the following command:

% tar xvf /cdrom/cdrom0/`more /images/imagelist.txt`

Do this for all the CD’s on the list.

###### Troubleshooting.

The reader reports that the CD-Archiver is not writing to the CDStudio

Log in as “Read” and type the following:

% rm cd_doing_now

A scanning site is reporting that the DICOM receiver is not working

Log in as “Read” and type the following:

% ps –cafe | grep simple

The last field of the output lists all the DICOM receiver port numbers which are running.  If the scanning site’s particular port number is not present, reboot the computer and restart the DICOM receiver.

###### XII. CT Reading Procedure Operations Manual

INTRODUCTION

When you first log on the workstation, the screen will look like the image in Fig 12.1.  The blue part of the screen is called the Workspace.  All the programs you will be using will be shown on the Workspace in separate windows like Windows98.  Moving around this operating system is just like, moving around Windows. You have a left mouse button to click on icons and commands.  To bring up various menus, use the right button.  Clicking the right button on any part of a window border will bring up various options.  These options will let you “CLOSE” the window, meaning minimize the window into a small icon, or “QUIT” programs and completely turn them off.

In the middle of the screen is a little gray box labeled “WorkSpace”. This is where you will launch various programs.  To bring up this “WorkSpace” box, you will click any part of the blue workspace using the third mouse button; this is known as “Right Clicking”.

There are many items on this workspace box that will launch various programs.  For now we will concentrate on the programs DICOM Viewer and Ca Tool Pipeline, which are the reviewing and scoring programs, respectively.

SUBJECT CHOOSER

##### 1.  Launching the Subject Chooser

First bring up the “WorkSpace” box, by right clicking the desktop.  Then you will right click the button labeled “Ca Tool Pipeline”.  After a few seconds, the two windows on the right (Fig 12.2.1) will appear.  The tall gray window on the left is the Subject Chooser menu of the Scoring Software.   This is where you will be choosing studies to Box, Trace and Score.

##### 2. The Subject Chooser Window

On the Subject Chooser window, you will find three main icons:

Boxing:	Used to launch the boxing program.  This is where you will outline the entire heart in a box.

Tracing: Used to launch the tracing program.  This is where you will identify the coronary arteries.

Scoring:  Used to identify plaque.

Under the icons are five buttons:

Update:     Refreshes the window, and updates the patient list

Rename:    Do Not Use

Delete:      Do Not Use

Rollback: Used to rollback studies from BOXED or TRACED status to NAMED. This will delete scoring and tracing files.

Abort:       Closes the Ca Tool program

Edit Scoring Defaults:  Setting defaults (see below)

On the right is a close up view of the subject list. This is where you will be choosing a study and left-clicking on it to highlight it.  The first sequence of numbers and letters is the study identifier.  To the right of this is the status of the study.   NAMED means that it has not been boxed yet.  BOXED means that it has been boxed and is ready to be traced.  TRACED means that it has been boxed and the coronary arteries have been identified but has yet to be scored.  SCORED means that it has been scored and saved.  To the right of the names is the scroll bar.  To scroll down, left-click on the bottom arrow.  Each click will move the patient list up and reveals a new name at the bottom of the window.

##### 3.  Edit Scoring Defaults

This is where you will set up the defaults for the Calcium Tool.  Fill out the fields (Fig 12.2.4) according to the criteria set by a particular study’s protocol.  You may alter the values either by directly entering numbers into the field and pressing ‘ENTER’ or by using the slider if one is provided.  This has to be done only once during the start of one session of reading.

Scans will be read using a HU threshold of >= 131, a minimum number of pixels of 4, a pad region of 8 pixels, a growth threshold of 101 HU and a search radius of 100 mm.

BOXING

##### 1.  Description of Boxing

#### Note that a reader must box every scan.

#### In the boxing window, you will view consecutive images of the patient’s heart.  You will then enclose the heart in a white box.  This procedure will produce the boxed view of the heart you will be using during the Tracing and Scoring procedures.

#### Another purpose of the boxing stage is to properly identify the hydroxy-apatite Phantom regions. To do this you will have to make sure that the bright white dots stay inside the cylindrical regions of the phantom.

#### To first get started you will have to launch the boxing program.  On the Subject Chooser highlight a patient study by first left-clicking on it.  Then move the mouse over the first icon labeled “Take Heart”, previously labeled in Fig 12.2.2.

#### When the boxing window pops up, it will cycle through all the slices of the patient’s images.  When this happens, you will have to keep an eye on the phantom to make sure that the bright white circles stay within the cylindrical regions of the phantom.  To check the slices again after the initial cycle, right-click anywhere on the boxing window (to bring up the menu) and choose “Cycle Through”.

##### 2.  Boxing the Heart

#### The next step in the boxing process is to actually box the heart.  This serves to isolate the image in the boxed portion for the next process, Tracing.  After the window finishes cycling through the images, there should be a white box surrounding the image of the heart. Resize the box by left clicking on either outline so the box completely encompasses the heart.

#### Once again you will have to check if all of the heart has been included in the boxed region.  Do this by right clicking to bring up the menu and then choosing “Cycle Through”.  Once you are sure the entire heart is boxed, bring up the menu once again and click on “Done: process”.  A small window will pop up on the top left corner of the screen. Choose “begin tracing” to launch the second part of the Scoring Software.

##### 3.  Identifying Regions in the Phantom

#### During the boxing stage you will be checking to see if the cylindrical regions of the calibration phantom have been properly identified.  Figure 12.3.3 shows a phantom whose cylindrical regions have not yet been identified.  And figure 12.3.4 shows a phantom with properly identified regions.

#### The purpose of identifying the hydroxy-apatite regions in the phantom is to let the scoring software have a way to compare the known density of the hydroxy-apatite phantoms to that of calcification in the arteries.

#### When you do come upon a region which has not been properly identified, right-click the center of the nearest hydroxy-apatite region and choose “set Phantom Region” from the menu (Fig 12.3.5).  The white circle nearest to the center of the hydroxy-apatite region you clicked will jump there. Do this with all the regions till the circles appear to be located within all regions.  Figures 12.3.5 to 12.3.7 shows this process in order.

TRACING

##### 1.  Launching the Tracing Program

There are two ways to launch the Tracing program: by clicking on the Trace button (see Fig 12.2.2) or by finishing the Boxing process as described in the previous section.  The purpose of the tracing program is to actually trace the path of the vessels through the heart, so that the Scoring Software will know where to look for lesions.  Basically you will scroll through all the slices and identify where the appropriate vessels are.

Navigating through the slices is simple.  First you must left-click the tracing window.  Then you move forward through the slices by pushing the ‘down’ arrow key on the keyboard.  To go in reverse, use the ‘up’ arrow key.

##### 2.  Labeling the Vessels

Now you will label the vessels.  Refer to Fig 12.4.2 to Fig. 12.4.4.  First scroll through the slices

using the arrow keys till you see the first appearance of the left main artery. In this particular case, it is on slice number six (Fig. 12.4.2). Now right click on the black space and choose “L. Main” from the menu that pops up.  A red “X” will appear over the vessel.

The computer will distinguish the vessels using colors: Red for Left Main, Green for the L.A.D., Yellow for the Circumflex, and Blue for the Right Coronary.

Next, look for any additional identifiable vessels on the same frame. If there are any, identify them the same way. If none, move on to the next frame.  In our case the LAD is also present, so we label it by right clicking on the vessel to bring up the pop-up menu, and then choosing “L.A.D.”. A yellow arrow will now appear.  On our next slice, Fig. 12.4.2, the L.A.D. is still present.  To label a vessel in the following frame just left click on the vessel and the program will label it accordingly.  For example, in Fig. 12.4.3, we left clicked on the LAD to continuously mark it and right clicked on the pop-up menu to label the circumflex and LAD branch.

##### 3.  Checking the Tracings

After you have labeled all the vessels in the slice where they last appear, you should check your work by right clicking anywhere on the image and choosing “Done: process” from the pop up menu. When you do this you will see the slices reappear consecutively, with the circular regions surrounding the points (Fig 12.4.5) where you placed the X’s.  Be sure that the entire coronary tree is included.

##### 4.  Editing the X’s

If you note that the regions are not encircling the arteries in every image, you should then choose “edit these results” from the little menu (Fig 12.4.6) you now see at the top of the screen.  Then page through the images and delete the incorrect points by using the ‘Shift’ key  and left clicking on the X’s you want to remove.  Then replace it with a new X in the correct position.  When you have finished editing the X’s, right click again, and choose “Done: Process” from the pop-up menu and return to section 3 above.

##### 5.  Finish Tracing

If the regions are placed correctly, choose either “Save and Score” from the resulting options (Fig 12.4.6) to proceed to the Scoring program or “Exit” to end the procedure and return to the Subject Chooser.

SCORING

##### 1.  Launching the Scoring Program

Launch the scoring program by either highlighting a ‘TRACED’ study in the Subject Chooser and clicking on the Scoring Icon (Fig 12.2.2), or by choosing “Save and Score” upon finishing the Tracing program.

##### 2.  Checking Lesions

If the scoring program cannot find calcium within the regions created during Tracing, the Quality Assurance Menu will automatically appear.

If the Program does find calcified lesions, you must decide whether or not these really represent coronary calcium.  In this case, the program will take you directly to its choice for a lesion.  These lesions (Fig 12.5.1) will be surrounded by a circle.  A circle of the same color will appear somewhere else on the image.  Move your mouse to drag the circle to a homogenous blood-pool near the lesion in question.  Click the left button if you agree that this is coronary calcium.  Otherwise click the middle button.  Then repeat this procedure for every “lesion” the program identifies in this image.  The program will automatically jump to the next image with a calcified lesion after you have assessed all the lesions in this image.

After you have assessed the last lesion in the last image, the QA menu will appear.

##### 3.  Quality Assessment Tool

The QA menu (Fig 12.5.2) has seven categories that you must assess.  These are Motion Artifact, Streak, Phantom, Misreg, Noise, Arteries, and Coverage.

Use the following guidelines:

Motion Artifact – Choose ‘Unacceptable’ if there is blurring of the distal RCA or other notable portion of the coronary arteries. Otherwise choose ‘Excellent’

Streak – Choose ‘Unacceptable’ if a streak artifact overlaps any calcified coronary lesion. Otherwise choose ‘Excellent’.

Phantom –  Choose ‘Unacceptable’ if there is insufficient assessment of the phantom in any image.  Otherwise choose ‘Excellent’.

Misreg –  Choose ‘Unacceptable’ if the images in the projection window do not show smooth epicardial borders.  Otherwise choose ‘Excellent’.

Noise –  Choose ‘Unacceptable’ if the image contains excessive noise that the program has mistakenly identified as coronary calcium.  Otherwise choose ‘Excellent’.

Arteries –  Choose ‘Unacceptable’ if part of a major coronary artery is cut at the edge of an image.  Otherwise choose ‘Excellent’.

Coverage –  Choose ‘Unacceptable’ if the origin of the left main coronary artery or the coronary sinus is not included in the image scan.  Otherwise choose  ‘Excellent’.

Click the “Save and Send” button after completing the quality assessment.

##### 4. Auto Archiving (Automatic CD Backup)

Immediately after you click the “Save and Send” button, the software automatically places a copy of the study you have just scored, including original images and all tracings, into a hidden archive folder.

After each complete scoring a window will pop up (Fig12.5.3), indicating the size of the archive.  Once the size of the archive reaches over 600 megabytes another pop-up window will appear asking if you would like to continue with the backup.  It will then ask you to confirm the writing of the backed-up archive files to the CD Studio.  Once you click continue in the window in Fig 12.5.4, you must then place your mouse over the “cmdtool” window of the Ca Tool software, Fig 12.5.5.

The software will then proceed to transfer the archive from the Sun Workstation onto the CD Studio.  If you continually check on the “cmdtool” window it will show the progress of the transfer which usually takes 15 minutes.  Once it has transferred the archived images, another window will pop-up asking if you would like to continue and write to the CD.  After you click continue, the archive will be burned into the CD.  This will take approximately 15 minutes. Please note that you may open another Ca Tool window and continue scoring while you are backing up.  Once writing is complete, the program will ask you to remove the CD and label it.

Test the CD to make sure all the files have been properly burned in.

APPENDIX A: Protocol for Inter-scan, Inter-reader, and Intra-reader Variability

This protocol is designed to serve as a guide for the assessment of inter-scan, inter-reader, and intra-reader variability.

Necessary log-in accounts will need to be created as follows:

For regular reading:   MESA-AP

For inter-reader and intra-reader variability:   MESA-RD, MESA-ER

##### 1.  Interscan Variability.  Since nearly all scans will be available in duplicate (except for those where one of the scans is technically inadequate for transmission or reading), the coordinating center will select the appropriate sample of studies for examining inter-scan variability among each of the scanning sites.  Transmitted summary of scores files will contain duplicate scans with the same identification number but different scan times.  The different scan time will be used by the coordinating center to distinguish between the first and the second scan.  Assessment may require as much as 25-50% of the duplicate studies to be examined in order to obtain a reliable monthly inter-scan variability for each center and/or technician.

2.  Inter-reader Variability.  5% of scans read and scored by the main reader will be reread and re-scored by a second reader on a quarterly basis.  Once the scans have been read and scores transmitted, the coordinating center will run a random number generator program that will determine which scans are to be re-read.  The selected scored scans will be loaded on the workstation and rolled back to “NAMED” status by the data manager.  The second reader will then be notified to read them.  After scoring, the reader will indicate on the software menu that this is a quality control study. Once read, the summary of scores file will be transmitted to the coordinating center where it will be evaluated for inter-reader variability.

3.  Intra-reader Variability.  The main reader will reread and re-score 5% of the scans on a quarterly basis.  Once the scans have been read and scores transmitted, the coordinating center will run a random number generator program that will determine which scans are to be re-read.  The selected scored scans will be loaded on the workstation and rolled back to “NAMED” status by the data manager.  The main reader will then be notified to read them again.  Once read, the summary of scores file will be transmitted to the coordinating center where it will be evaluated for intra-reader variability.

APPENDIX B: Description of Options in Scoring Defaults Menu

Introduction.  The tracing program creates a tubular region of interest of radius R.  These tubes define the “voxel list”.  Only these tubes of diameter R are available for evaluation of calcified plaque.  R has been set at 7mm.  This value is a compromise between higher values which would necessitate a lot of pruning and lower ones which would necessitate an excessive amount of editing during tracing.

1.	Source.  This can have two values: Voxel List or Pixels.  Always set this as Voxel List.  Setting at Pixels will cause the tracing program to search outside of the voxel list.

2.	Edits.  This can have three values:  Override, Show, or Ignore.  These are important in rescoring a case and in batch scoring.

Override – causes the scoring program to override prior edits and thus to start “from scratch”.  This should always be used in interactive scoring.

Show – causes the lesions which have been edited out or removed during a prior scoring to be listed but not to be counted in adding up the score.  They are listed in an intermediary file (REG) as noise.

Ignore – causes the scoring program to include edits from prior scorings and therefore to ignore regions which were removed during prior scorings.  Ignore should usually be used for batch scorings.

3.	Extract.  This allows the user to choose whether he/she wants to use the actual pixel values (Brightness in Hounsfeld Units) or to scale these to the phantom (Scale to Phantom).

4.	Skip Selected Levels.   Do not use this option.

5.	Apply Proximity Limit.  Possible values for this are Nearest, Farthest, Peak and Each.

Nearest – causes plaques which are on the border defined by the search radius, R, to be entirely included.

Farthest – causes such plaques to be entirely excluded.

Peak – causes these plaques to be included only if the peak value is inside of R.

Each – causes the plaque to be split so that only the part within the R is included.

In general, always use Nearest.

6.	Include calcium at base.  This can be True or False.  True causes ostial calcium to be included.  False causes it to be excluded.  Set this at True.

7.	Reference Region Filter.  Set as Brighter than Threshold.

8.	Pad Region.  This defines, in units of pixels diameters, the distance the scoring program will search from a calcified lesion defined by the base value and minimum lesion size.  This will affect the arterial summation calcium mass estimates but should never affect the other measures.  Set this at 8 pixels for interactive scoring.  Set this at maximum for batch scoring.

9.	Growth Threshold.  This defines in HU the pixels included in the arterial summation mass formula.  All pixels of at least this value within the pad region will be included in the arterial summation mass formula.  Set this at 101 HU for interactive scoring.  Set this at maximum for batch scoring.

10.	Vessels.  This allows the user to exclude or include vessels in scoring.  By blanking out a vessel, the reader causes a zero contribution from that vessel to be counted in the summary of scores file.  The vessels allowed are: Aorta, Left, Main, LAD, CX, RCA, Noise, Other.   In general only Left Main, LAD, CX, and RCA should be highlighted.

11.	Search Radius [r].  This, in mm/10, defines the outer limit of the search for lesions and the outer limit for the extension of lesions for the arterial summation calcium mass algorithm.  It is limited by the radius of the voxel list, R.  Therefore, if r is greater than R, R defines the search limits.  Set the search radius at 100 mm/10 for both interactive and batch scoring.

12.	Base Value.  This defines a calcified plaque.    Set it at 131 HU for both interactive and batch scoring.  Do no change the base value.

13.	Minimum Lesion Size.  This defines the minimum size of a calcified plaque.  It pertains to the minimum number of contiguous pixels (sharing a side) that will define a calcified lesion.

APPENDIX C:      IDENTIFICATION AND REPORTING OF SCANNER PROTOCOL VIOLATIONS

CT Reading Center Identification and Reporting of Protocol Violations.

The following procedure will be followed regarding scanner protocol violations.

1.   A clear protocol will be made available to the scanning centers at the initiation of the study and whenever there is a scanner or protocol change dictated by the director of the Reading Center.  This will be personally sent to the lead technologist AND discussed at a site visit or during a documented phone conversation between the reading center director or co-director and the chief technologist before the first participant is scanned.  It is the responsibility of the director or co-director to insure that the technical staff use and follow the protocol.

2.   The reading center director or co-director will edit the dicom header checker text file to include the relevant elements of the new protocol.  The director or co-director will copy the elements of this revision directly from the MOO or written protocol that has been sent to the technologists. This chore of editing the dicom header checker file will be done personally by the director or co-director and will not be assigned to another employee.  The systems administrator will always be witness to this change when it is made, not later or when it is convenient.  The systems administrator will approve all such changes in the dicom header checker file.

- After changing the dicom header checker file, the director or co-director will send an email to the director, the data manager, systems administrator and the reader.  This email will state  that a revision of the dicom header checker file has been made, exactly what the revision is (all elements, mA, etc) and instructing the reader and data manager to watch for protocol deviations.

- Variations of no more than 10% will be permitted for all continuous protocol parameters (mA, KVp, exposure time, pixel length, slice thickness, etc)

- The director or co-director in the presence of the systems administrator  will test the revised dicom header checker file by reading a scan known to have been done in protocol and, when possible, one known to have been done out of protocol to be sure that the software is functioning properly.

- When the reading software detects a scan for which the dicom header does not match the dicom header checker file, the reader will complete a form indicating study id number, date and time, site, study.

- The reader will give this form to the systems administrator and the data manager. The data manager will inform the scanning site involved, the PI of the study and the local PI at the clinical center by email of the nature of this deviation. This email will be copied to the systems administrator, director and co-director and when applicable to the coordinating center for the study.

- The systems administrator will carefully read all such emails (7 above) and  match them with the forms that he has been given by the reader (6 above).

APPENDIX D: CD Backup Procedure (Manual)

1)	Place a blank CDR into Young Minds device.

2)  	Select studies to backup and place in a separate folder.

3)  	Left click the folder, containing the files to be backed up.

Then right click “File” ( “Custom Commands” ( “How Big?” to determine the size of the folder.  The value that is reported is in kilobytes and it should be less than 640,000.

4)	When you are ready to backup,

Right click “File” ( “Custom Commands” (”MakeDisc”.

5) 	The MakeDisc program window will appear.

First, Left click the “Base Directory” button to specify which folder contains the data to be backed-up.  In the ‘selections’ space, type the name of the folder (i.e. /home/rd/forbackup).  Click “O.K.”

Second, left click “image” ( “Premaster”, to start the backup process.

Once dumping has been completed, you may then “left click” “Image” ( “Cut” to begin writing to the CDR.

6)	Once the status bar has indicated that the disc is finished, eject the disc, label it, and close the MakeDisc program.

APPENDIX E: How to Make and Test A Test Case

The test case software has been created by Mr. Reed in order to allow the user to check up on the software periodically, to ensure that it is working correctly.  All readers should know how to make and test a test case.

- First, delete the “testcase” already present by using the cmdtool (command tool).

%rm  –ri  /images/testcase

2) 	Select a good case (one with a cal. phantom and clear arteries).

3) 	Box it using Reed Scoring software.

4)	Copy its directory to the “testcase” directory using the cmdtool (command tool).

% cp  –r  /images/(name of study)  /images/testcase

5) 	Click the “Update” button on the subject chooser and trace the “testcase” you just made.  Place X’s only where you would like to “create” a known calcified lesion.

After Tracing, choose “DONE: Process” ( “Save and score”. Then quit the scoring window.

6)	Using the ‘File Manager’, rename the file “trajectory” with the name “phanplaque”.

Right-click desktop to bring up the “Workspace menu”

Right-click “Programs” ( “File Manager”

Use the “File Manager” to bring up the ‘/images/testcase’ directory.

Left-click the file ‘trajectory’ file (it will be highlighted).

Type in ‘phanplaque’ ( press “Enter”.

# Create calcified lesions.  You can create one or more than one.  For starters, try a simple case like the following:

| 150 | 140 | 131 |
| --- | --- | --- |
| 90 | 240 | 156 |
| 140 | 100 | 166 |

# Assume that the default Hounsfeld number defining a lesion is greater than 130 and that the default minimal lesion size is 3 adjacent pixels (.8 mm2 with a 30 cm F o V). Remember that adjacency is defined as sharing a side, not a corner. So the score for the above lesion would be (number of adjacent pixels over 130) x (weight determined by max HU in lesion) x (area of one pixel).  For a 30 cm fov and a minimum lesion size of 3, this gives a score of: 6 * 2 * (.52 * .52) or 3.24.

# Now to make this lesion in the ‘phanplaque’ file, double-click the file in the “File Manager” and edit as follows.

# The lesion making program only makes square lesions like the one above. Delete the first two numbers from the first line of the file.

# Add a lesion size, in pixels, ‘n’ at the end of the first line.

# Press “enter” to go to next line.

# Add matrix of ‘n’ lines and ‘n’ columns of densities of “HU” numbers.  In above case, this would be 3 columns and 3 lines.  Separate these entries by tabs.

Repeat this process for as many lines in the original file as you like.

Save by Right-clicking “File” ( “Save”

Record on a piece of paper the matrix for the lesion you have created.  Score it by hand and record the score under your matrix.

Exit Text Editor by right clicking “File” ( “Quit”

8)	Run the program ‘fake_phan’ in the ‘testcase’ directory by using the cmdtool

% cd  /images/testcase

% fake_phan

9)	Save results in ‘boxed’ image file.

% mv  the_phantom  the_heart.skw

- Trace and score the case called “testcase” in the subject chooser.

- Check the results and compare them with what you calculated in number 7 above.

- When you have created the test case correctly.  Delete the ‘.img’ file from the test case directory by typing in the following:

% rm –i /images/testcase/*.img

XIII. ATTACHMENT

Attachment 1

###### Attachment 2

CT Study Transmission Sheet/ FAX to 310-533-1616

CT STUDY TRANSMISSION SHEET

STUDY (CIRCLE):       	MESA

Transmission	Date ______/ _______ / __________

Transmission 	Time _________:__________

Scanning Center Location: ______________________  	ID: _________________________

Staff sending form: ____________________________	Phone: ______________________

| ID # | Acrostic | Date Scanned | Archived (Y/N) | Number of Slices | Comments |
| --- | --- | --- | --- | --- | --- |

Attachment 3

QA QCT Sheet - FAX to 310-533-1616

Quality Assurance QCT Data Sheet for MESA Study

Mesa Site _____________________________ CT Scanner ____________________________ Mfgr._________________________________

CT Scanner Model______________________ Phantom Serial Number_________________

| QA Date | Tech. Initials | Kvp | MA | FOV | Slice Thickness | Table Height | Hounsfeld Numbers |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | 0 | 50 | 100 | 200 | Torso PIN |

Attachment 4

MESA CT Reading Center Complete Scan Log

| ID # | Scan Date | Scan Time | Send Date | Check Sum | Send Time | Checked for completeness | Phantom |
| --- | --- | --- | --- | --- | --- | --- | --- |

Attachment 5

Attachment 6

CT Technician Certification Checklist for Performing Coronary CT Scans

HARBOR-UCLA CT Reading Center

Please complete and fax to Dr. Nathan Wong at 310-533-1616 or email ndwong@uci.edu

Study (circle):  MESA   Other (specify):___________________________

Site / Institution: ____________________

Name of Technician to be certified: _____________________

Phone: ___________________ Fax: ___________________

E-mail: __________________

Evaluator (Currently Certified Technician) (please check as participant completes):

___ Professional greeting of patient

___ Correct placement of calibration pad phantom

___ Teaching of breathholding technique

___ Selection of correct field of view (35 cm) including phantom

___ Scanning of entire heart (at least 35 slices)

___ Appropriate lead placement and ECG gating

___ Instruction on transmission of images via internet

___ Instruction on proper ID labeling of study in header

___ Instruction on QC procedure for CT Calibration using Calibration and Torso phantoms

Signature of Evaluator / Investigator Pre-Certifying Technician:          Date:

___________________________________________                           ____________

NOTE THREE SCANS OF SATISFACTORY QUALITY DONE ACCORDING TO PROTOCOL MUST BE TRANSMITTED TO THE CT READING CENTER PRIOR TO RECEIVING FULL CERTIFICATION, WHICH WILL BE CONFIRMED VIA E-MAIL.   Please list the IDs, acrostics, and scan dates of these three scans for certification here (and on the usual fax transmission sheet):

| Patient ID | Acrostic | Date Scanned |
| --- | --- | --- |

FOR CT READING CENTER USE:

_____ Confirmation of Receipt of 3 Scans of Suitable Quality by CT Reading Center

_______________________________________________________________________

Signature of CT Reading Center Investigator             Date

Additional notes/problems:

# Text boxes

Fig  12.3.5

Fig. 12.3.5

Fig 12.2.4

Fig 12.3.3

Fig 12.1

Fig 12.4.6

Circle that follows the mouse.  This must be in a homogenous blood pool when assessing a lesion

Fig 12.3.6

Calcified lesion which program has identified.

Fig 12.5.1

Fig 12.5.3

Fig 12.5.4

Torso

Fig 12.5.2

Fig 12.5.5

isocenter

9.2 cm

Cal. Phantom

L.A.D

Table

Fig 12.1

Fig 12.4.2

Slice Number

Fig 11.4

Left Main

L.A.D.

Fig 12.4.3

Right-Click Menu

Fig 11.2

Fig 11.1

Fig 11.3

Fig 11.5

Fig 11.6

Changes directory to /images

This command makes a directory.  I suggest you use the study name, there will be less confusion if there multiple studies on one workstation.

The Target Directory.  Change “studyname” to whatever directory you created in the previous command.

The folders to be moved.  Replace “yyyymmdd” with the appropriate date, i.e. “20010301” will move all studies scanned on March 1, 2001.

Renames the file.

This means to delete all folders in “/images” that have are identical to any entries in the file “mesalistxxx”.

Unix delete command

Extra Info;

The files on the CD are TAR files; these are similar to ZIP files on PCs.  One TAR file actually contains the whole study folder (including images and tracings).  It is also labeled with the same name as its original study folder in the “/images” directory, plus the date the TAR file was created.

Fig 11.8

Fig 11.7

Excel Files:

For Cardia and Mesa, the excel Log Sheets are located in Chris Dailing’s computer:

/MyDocument/Chris/EBCTSCDS/Daily Log Sheets/…

Fig 12.3.7

Circumflex

Fig 12.4.4

Right Coronary

Fig 12.3.6

Fig 12.3.1

Phantom

Fig 12.2.2

Boxing

Tracing

Scoring

Right-Click Menu

Fig 12.2.3

Scroll Bar

Fig 12.2.1

Ca Tool Scoring Software

Fig 12.3.2

White box outlining heart

Surface Rendered Region

Cylindrical Hydroxy-apatite region

Fig 12.3.3

Properly centered regions

Fig 12.3.4

Fig 12.3.5

Right click menu.

Fig 12.3.6

Circular region surrounding a vessel previously marked with an X.

Fig 12.4.5