## Slide 1: Exam 6 Seated Blood Pressure
Cathy Nunn
Wake Forest University
Study Coordinator

## Slide 2: Seated Blood Pressure
Measurements in the MESA Study:
Arm Circumference
Systolic Blood Pressure
Diastolic Blood Pressure
Pulse rate
Pulse Oximetry

## Slide 3: Contraindications
Avoid taking measurements on right upper arm if participant is affected on the right side by:
stroke,
mastectomy,
renal fistula, or
if the participant has IV catheter

## Slide 4: Safety
Alert levels requiring immediate referral (send participant directly to a physician or hospital) for MESA participants are:
 Systolic BP >210 mm Hg
 Diastolic BP >120 mm Hg
Alert levels requiring urgent referral (within one week) are:
 Systolic BP 180–210 mm Hg
 Diastolic BP 110–120 mm Hg
Alert levels requiring follow-up within two months time, and, therefore, we recommend physician notification for systolic or diastolic BP above these levels.
 BP >140/90 mm Hg

## Slide 5: Equipment
Blood pressure and heart rate measures
 Dinamap Monitor Pro 100 automated blood pressure device
 Blood pressure cuffs (Dura-cuf Adult Assortment Pack.
 Measuring tape (for arm circumference).
 Watch or stop watch (to time rest period and resting heart rate).
 Hand calculator.
 Copy of Critikon chart for choosing correct BP cuff size.
 Information sheet on interpretation of BP.
 Resting Heart Rate/Blood Pressure Form.
Pulse oximetry measures
 Devon Medical Handheld Pulse Oximeter PC-66

## Slide 6: Blood Pressure Measurement
Setting up the Dinamap BP Machine
Load the printer paper by opening the flap on the side of the device.
Turn on the Dinamap device, push the "Off/On" button on the front control panel (lower left).
After five seconds an initial message will appear on the LCD screen. It will consist of a warning and the instruction, “push a front panel key to start.”
In the main menu select print using the gray toggle knob. In the next menu, select auto and then push the toggle knob. This will program the device to print the blood pressure measurements.
Do not touch the monitor again until you have completed cuff placement, rest period, and and you are ready to proceed with blood pressure measurement.

## Slide 7: Blood Pressure Measurement
Positioning the Participant
The workstation should be free of excessive noise or distractions.
The participant should be seated and relaxed in a comfortable chair:
 He or she is sitting up (not slouched).
 Both feet are on the floor (legs/ankles not crossed).
 Right forearm is supported resting on the table.
The participant should not talk, eat, or drink during the procedure.
Dinamap output should not be visible to the participant during the measurement.

## Slide 8: Blood Pressure Measurement
Application of the Blood Pressure Cuff
Place the appropriate cuff around the upper right arm so that the mid-height of the cuff is at heart level. Palpate the patient’s brachial artery and place cuff so that the artery is aligned with the cuff arrow marked “artery.”
Place the lower edge of the cuff, with its tubing connections, two centimeters above the natural crease across the inner aspect of the elbow.
Wrap the cuff snugly around the arm, with the palm of the participant's hand turned upward.
Secure the wrapped cuff firmly by applying pressure to the locking fabric fastener over the area where it is applied to the cuff.
Do not wrap the cuff too tightly around the arm. You should be able to insert the first joint of two fingers under the cuff. The cuff should be snug but not tight.

## Slide 9: Blood Pressure Measurement
Rest Period
The participant should rest for five minutes (timed using a watch or stop watch) prior to the heart rate and blood pressure measurement.
When the five-minute rest period is over, but before the first blood pressure measurement is started, record the time of day on the Seated Blood Pressure Form (examples: 04:25 P [p.m.] or 11:38 A [a.m.]).

## Slide 10: Blood Pressure Measurement
Blood Pressure Measurement
To begin the blood pressure procedure, access the Main Menu on the Dinamap by pushing the “Start/Stop” button at the lower right of the monitor. (Please note that the gray knob located at the upper right of the monitor allows you to change selections in the monitor screen, in a manner similar to a computer mouse or pointing device. Rotate the knob in order to move from one item to another in the monitor screen, and push it to select the desired option.)
Use the knob to select the set bp option from the menu and then press the knob (equivalent to clicking a mouse) to implement the selection. The next menu appears automatically.
Use the knob to select auto bp and then press the knob.
Immediately after you select auto bp the monitor will start the first blood pressure measurement. However, during this first inflation, select the window that has appeared to the right of auto bp and push the knob, so that there is a black number against a clear background in the window. Rotate the knob to select “2.” This will select two minutes as the interval between sequential blood pressure measurements. Push the knob again (colors in window will reverse) to implement the selection. (The device will retain this setting, even after it is turned off, so you will not have to repeat this step again.)

## Slide 11: Blood Pressure Measurement
Palpate the radial pulse during inflation. The radial pulse should not be palpable at peak inflation pressure. If the participant's radial pressure remains palpable when the device begins to deflate, the device will complete its deflation procedure and then should automatically reset itself for a higher inflation pressure and repeat the measurement.
In the unlikely event that higher inflation pressure is needed but the machine does not automatically reset to a higher pressure:
Check carefully to be sure that the cuff is properly positioned on the participant's arm with the arrow at the brachial artery.
manually reset the inflation pressure (210, 260, 300).
It is not necessary to repeat or prolong the five-minute rest period, if this happens, but explain the change in the procedure to the participant (e.g., “I think we need to use a higher inflation pressure—I'm just going to reset the machine”).

## Slide 12: Blood Pressure Measurement
When the radial pulse is obliterated at maximal inflation, the first blood pressure measurement will be obtained.
The device will automatically obtain the 2nd and 3rd measurements, at two-minute intervals.
Record the three sequential blood pressure (and heart rate) readings on the Seated Blood Pressure Form.
Print the results and set the machine to manual.
Remove the blood pressure cuff from the participant's arm and thank the participant for his/her time.

## Slide 13: Blood Pressure Measurement
BP Measurement Instructions for Participants With Short, Thick Arms
Occasionally there will be a participant whose upper arm is too thick and short for the thigh cuff or on whom the thigh cuff pops open on inflation. The alternative procedure in this case is to obtain the resting blood pressure in the right forearm.
Measure the forearm circumference at the midpoint between the olecranon (elbow) and the ulnar stylus (wrist bone on pinkie side). Select the proper size cuff based on the forearm measurement. The blood pressure procedure is otherwise the same.
You must document on the Seated Blood Pressure Form that you have measured the forearm blood pressure.

## Slide 14: Reporting results
Reporting Blood Pressure Results to Participants
The technician may verbally provide the participant with the blood pressure reading (the average of the last two pressures), if asked, after the procedure has been completed.
If the blood pressure is normal (<140/90), the technician may say that it is normal, particularly if asked.
If the blood pressure is not normal (>140/90) but not at an alert level (>210 mm Hg), the technician should exercise the standard option of not discussing the interpretation or stating that it does appear to be high (or “somewhat elevated”) but that, again, it will be discussed later.
If an alert level is identified, the technician should calmly notify the clinic coordinator when the procedure has been completed. (If symptoms of severe hypertension are present, the technician should notify the clinic coordinator immediately.)

## Slide 15: Maintenance and calibration
Once a week each device should be used simultaneously with a paired device to simultaneously measure the blood pressure in each arm of a non-smoker under the age of 50, in whom there is no reason to suspect that the blood pressure in the two arms should differ.
Repeat the measurement three times.
If the paired blood pressure measurements agree within 4 mm Hg or less, for both systolic and diastolic BP, the devices are considered to be in calibration.
Investigate any systematic divergence, even if less than 4 mm Hg (e.g., by switching arms and/or pairing the devices with a third device).
If the two devices differ by more than 4 mm Hg, calibration must be done. It should be recognized that, if the cuff deflation rate is 2 mm Hg/sec and the heart rate is 60 bpm, divergences of 2–4 mm Hg would be expected, even if the device is in perfect calibration.

## Slide 16: Pulse Oximetry Measurement
Preparation
Resting oxygen saturation will be measured while the participant is resting for blood pressure measurement.
Explain the procedure to the patient
Verify that the probe is clean, dry and in good condition before applying it to the participant
Ask the participant to remove her nail varnish or acrylic nails that may impair the effective transmission of light (have some nail varnish remover on hand just in case).

## Slide 17: Pulse Oximetry Measurement
Positioning the Participant
The workstation should be free of excessive noise or distractions
The participant should be seated and relaxed in a comfortable chair
The participant should not talk, eat, or drink during the procedure

## Slide 18: Pulse Oximetry Measurement
Reading the pulse oximeter
Place oximeter on participant’s finger during the beginning of the resting period prior to blood pressure measurements.
Verify that the probe is well positioned
Record the apparent median value obtained while observing the monitor over a one-minute observation period

## Slide 19: Pulse Oximetry Measurement
For participants using oxygen
If the subject is using oxygen, the pulse oximeter should be placed on the subject‘s finger first.
Next, the subject‘s oxygen should be discontinued while monitoring the oximeter for a period of five minutes.
  If the pulse oximeter reading falls to 82% or less, oxygen will be replaced and a reading of 82% will be recorded as the subject‘s oximetry.
  The apparent median value obtained while observing the monitor over a one-minute observation period should be recorded.

## Slide 20: Recording Values
Enter values carefully – once the participant leaves, the numbers cannot be checked
The software will check values for plausibility and alert you of questionable values
To continue after an implausible value is detected, you can either confirm or correct the value

## Slide 21: Maintenance
Clean sensor/probe as per manufacturers‘ guidelines in between each patient use.
Calibration as per manufactures instructions
Daily quality control check on self
Carry spare alkaline batteries
To avoid any leakage from the batteries, remove them if the pulse oximeter is not in regular use
Report immediately to the Clinic Coordinator if the pulse oximeter appears to be malfunctioning