TrueMass® Total Fatty Acid Panel

Validation Report

TrueMass® Total Fatty Acid Panel

The TrueMass Total Fatty Acid Panel quantifies 38 individual fatty acid metabolites involved in lipid metabolism. The concentration and complete fatty acid composition are determined producing a high-resolution profile of basic lipid metabolism. TrueMass Total Fatty Acid Panel quantifies the lipid metabolites from almost any biological matrix including plasma, serum, tissues and cells.

List of Fatty Acids: 14:0, 15:0, 16:0, 18:0, 20:0, 22:0, 24:0, 14:1n5, 16:1n7, 18:1n7, 18:1n9, 20:1n9, 20:3n9, 22:1n9, 24:1n9, 18:2n6, 18:3n6, 20:2n6, 20:3n6, 20:4n6, 22:2n6, 22:4n6, 22:5n6, 18:3n3, 18:4n3, 20:3n3, 20:4n3, 20:5n3, 22:5n3, 22:6n3, 24:6n3, plasmalogen derivatives of 16:0, 18:0, 18:1n7 and 18:1n9, t16:1n7 t18:1n9 t18:2n6.

Validation Experiment Design

In order to evaluate the performance of our TrueMass Total Fatty Acid Panel, Tethys processed a validation experiment that included the analysis human serum samples from ten different subjects, run on 2 days days, each sample run in quintiplicate per day.

Assay Methodology

The lipids from serum/plasma and tissues are extracted in the presence of authentic internal standards by the method of Folch et al (Folch, J., M. Lees, and G. H. Sloane-Stanley. 1957. A simple method for the isolation and purification of total lipids from animal tissues. J. Biol. Chem. 226: 497–509), using chloroform:methanol (2:1 v/v). The total lipid extract is trans-esterified in 1% sulfuric acid in methanol in a sealed vial under a nitrogen atmosphere at 100°C for 45 min. The resulting extract is neutralized with 6% potassium carbonate and the fatty acid methyl esters (FAME) are extracted from the mixture with hexane.  The isolated FAME extract is then prepared for gas chromatography by sealing the hexane extracts under nitrogen. Fatty acid methyl esters are separated and quantified by capillary gas chromatography (Agilent Technologies model 6890) equipped with a 30 m DB-88MS capillary column (Agilent Technologies) and a flame-ionization detector.

Quality Assurance/Quality Control

Tethys employs a multi-layered QA/QC system to ensure the delivery of high-quality data.  Tethys has established a comprehensive quality assurance program that monitors and evaluates the ongoing and overall quality of the testing process.  Quality assurance begins at the time of sample collection.  Tethys provides recommendations for sample collection, handling and transport (contact Tethys directly for these recommendations).  Quality assurance continues throughout the analysis and data handling process to monitor all aspects of the process.

Tethys has established a detailed quality control system to ensure high quality data.  Tethys has a written procedure manual for the performance of all analytical methods used by the laboratory.  These procedures include performance of the procedure, calibration and control procedures, rejection criteria, and limitations for the method.  The laboratory has established and verified performance specifications, including accuracy, analytical sensitivity, and the reportable and the reference range of results.

All equipment maintenance and function checks are conducted as specified by the manufacturer and are documented.  Prior to any analysis, each instrument performs qualification experiments to determine if the instrument is capable of generating quantitative measurements, including analysis of reference standards.

Samples are analyzed in batches, which are defined as a group of samples that are analyzed in series within a day.  Sample quality is assessed on a per-batch (standard sample and LOQ analysis) or per-sample (internal standard, LOD and biological analyses) basis.  Quality control samples are included with each batch to continuously monitor the accuracy of the platform with time.  Each batch must meet defined acceptance criteria, including minimum and maximum thresholds of each analyte and the surrogate, quantitative and compositional precision and accuracy of the control samples, and pre-set Alert and Action limits.  In addition, data quality is assessed by comparing the values generated to values with the same or similar traits (matrix, species, etc) within our database. The laboratory routinely performs control procedures to monitor the stability of the systems utilized by the laboratory, and both accuracy and precision are monitored over time.

Imprecision

For calculations of imprecision, the serum samples obtained were prepared and analyzed in quintuplicate over several consecutive days.  Within-day imprecision was calculated as average percent coefficient of variation (%CV) for each consecutive day.  The between-day imprecision was calculated as the %CV of the mean values from each of the runs.  The concentration of the metabolites measured in this serum sample set is representative of most matrices.  The following tables report the quantitative and compositional within-run and between-day imprecision found in the serum matrix.

Quantitative and Compositional Imprecision

Quantitative imprecision was determined by calculating the between-day and within-day imprecision of TrueMass when expressed in nMoles per gram.  Compositional imprecision was determined by calculated the between-day and within-day imprecision of TrueMass when individual metabolites were expressed as a mole percentage of the sum of fatty acid metabolites. Absolute concentration and absolute mol percentage concentration, between-day imprecision  [%CV (Between Day)] and within day imprecision [%CV (Within Day)], values obtained are shown in Table 1.

Concentration Dependence of Imprecision

The imprecision of individual analytes is strongly influenced by their concentration in test matrix.  Thus, although the test serum samples used in the validation were a useful approximation of TrueMass imprecision for unknown plasma or serum samples, individual samples with different lipid compositions will demonstrate slightly different imprecision values.  The relationship of between-day and within-day imprecision to concentration is specific.  Values for this plot was obtained from the imprecision of concentration of each fatty acid analyte in the present imprecision experiment Figure 1.   As shown in this figure fatty acids greater than 0.10 mole percent have of % CV less than 20%.

TABLE 1

Figure 1